🧠 Abstract – IV Fluid Monitoring and Alert System using ESP32

Objective

The aim of this project is to design and develop an IoT-based Intravenous (IV) fluid monitoring system using an ESP32 microcontroller, HX711 load cell amplifier, and 16x2 I²C LCD display.
The system continuously monitors the IV bottle’s fluid level by measuring its weight, provides real-time display, and sends Telegram notifications automatically when the fluid level drops below pre-set thresholds — ensuring timely nurse or doctor intervention.

Working Methodology

1. Initialization & Setup

- The ESP32 initializes peripherals such as Wi-Fi, HX711 weight sensor, LCD display, and push buttons.

- The system connects to a predefined Wi-Fi network to enable Telegram message communication.

- The LCD display shows status messages like “Connecting WiFi,” “Place IV Bottle,” and “Waiting…” during startup.

2. Bottle Detection & Calibration

- Once powered, the system waits for the IV bottle to be placed on the load cell.

- The HX711 sensor measures the stable weight of the bottle.

- If the weight is below 220 g, the system identifies it as an insufficient or missing bottle and displays “⚠ Bottle too light!” until a proper one is placed.

- When the valid bottle is detected (above 220 g), its weight is stored as the reference (bottleWeight).

3. Threshold Setting (User Input)

- After bottle stabilization, the user is prompted on the LCD to set a custom IV threshold level (in milliliters).

- Three push buttons are used to set the threshold in hundreds, tens, and units:

  - Button 1 → increases the hundreds digit

  - Button 2 → increases the tens digit

  - Button 3 → increases the unit’s digit

- The selected value is displayed live on the LCD (e.g., “Thr: 245 ml”).

- If the user does not provide input within 20 seconds, a default threshold (e.g., 70 ml) is automatically set.

Circuit Pinouts

| Purpose / Component | Defined Pin Name | ESP32 GPIO Pin | Connection Description |
| --- | --- | --- | --- |
| HX711 – Data (DOUT) | #define DOUT 4 | GPIO4 | Connect HX711 DT pin here |
| HX711 – Clock (SCK) | #define SCK 5 | GPIO5 | Connect HX711 SCK pin here |
| LED (alert) | #define LED_PIN 25 | GPIO25 | Status LED or alert indicator |
| Buzzer | #define BUZZER_PIN 26 | GPIO26 | Active buzzer or alert buzzer |
| Button – Hundreds | #define BTN_HUNDREDS 32 | GPIO32 | Sets 100s place (connect button to GND) |
| Button – Tens | #define BTN_TENS 33 | GPIO33 | Sets 10s place (connect button to GND) |
| Button – Units | #define BTN_UNITS 27 | GPIO27 | Sets 1s place (connect button to GND) |
| LCD 16x2 (I2C) | (fixed internally) | SDA = GPIO21 / SCL = GPIO22 | LCD I2C module (0x27) |
| Wi-Fi / Telegram | — | — | Uses default ESP32 Wi-Fi hardware |

| HX711 Pin | Connects To |
| --- | --- |
| VCC | 3.3V |
| GND | GND |
| DT | GPIO4 |
| SCK | GPIO5 |

| LCD Pin | Connects To |
| --- | --- |
| VCC | 5V |
| GND | GND |
| SDA | GPIO21 |
| SCL | GPIO22 |

| Button | Connects To | Connect Other Side To |
| --- | --- | --- |
| Button 1 (Hundreds) | GPIO32 | GND |
| Button 2 (Tens) | GPIO33 | GND |
| Button 3 (Units) | GPIO27 | GND |

| Component | Positive Pin | Negative Pin |
| --- | --- | --- |
| LED | GPIO25 (via 220Ω resistor) | GND |
| Buzzer | GPIO26 | GND |

4. Continuous Monitoring

- The ESP32 enters the monitoring mode, reading weight values continuously from the HX711 every few seconds.

- The current weight and threshold value are displayed on the LCD in real time.

- The system calculates the amount of fluid used as the difference between the initial bottle weight and the current weight.

5. Smart Notifications & Alerts

- The total threshold range is divided into three parts (25%, 50%, 75%) to ensure early warnings:

  - 1st Alert (25%) → Early notification

  - 2nd Alert (50%) → Mid-level warning

  - 3rd & Final Alert (75%) → Urgent refill required

- Each alert triggers:

  - LED and buzzer blinking

  - Automatic Telegram message sent to the medical staff

- The final alert also includes a confirmation message like
“✅ IV Threshold completed for [Patient Name]”.

6. Safety & Fault Handling

- If the bottle is removed or the sensor reads zero weight, the system pauses monitoring and displays
“⚠ No Bottle Detected. Waiting…”

- Once the bottle is replaced, monitoring resumes automatically.

- Any invalid threshold (greater than the bottle’s capacity) is rejected, prompting the user to re-enter a valid value.

7. Communication & Cloud

- The Telegram Bot API is used to send real-time notifications over the internet.

- This approach allows nurses or hospital staff to monitor multiple patients remotely through a mobile phone or PC.

Advantages

- Prevents IV fluid overflow or emptying.

- Reduces manual supervision by nurses.

- Provides real-time monitoring and alerting.

- Uses low-cost hardware (ESP32, HX711, load cell, LCD).

- Can be expanded for multi-patient monitoring or cloud-based hospital dashboards.

Conclusion

The developed system ensures accurate IV fluid level measurement using a calibrated weight sensor and provides automated smart alerts via Telegram and display feedback.
This enhances patient safety and nursing efficiency, making it a reliable, affordable, and scalable IoT-based healthcare solution.